REAGEN™

Aflatoxin B1 ELISA Test Kit

RNM98007

### Product Description

REAGEN™ Aflatoxin B1 ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Aflatoxin B1 in cereals, meat/fish, feed, milk ,oils, peanuts and pistachios .The unique features of the kit are:

- Rapid extraction method .

- High sensitivity (0.05 ng/g or ppb) and low detection limit for milk (0.1 ppb), feed (1.0 ppb), and beef， pork, or poultry (1.0 ppb)

- High reproducibility.

- A quick ELISA assay (15 minutes incubation time).

### Procedure Overview

REAGEN™ The method is based on a competitive colorimetric ELISA assay. The Aflatoxin B1 antibody has been coated in the plate wells. During the analysis, sample and Aflatoxin B1-horseradish peroxidase (Aflatoxin B1-HRP) conjugate are added to the wells for incubation. If the Aflatoxin B1 residue is present in the sample, it will compete for the Aflatoxin B1 antibody, thereby preventing the Aflatoxin B1-HRP from binding to the antibody attached to the well. The resulting color intensity, after addition of the HRP substrate (TMB), has an inverse relationship with the Aflatoxin B1 residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Aflatoxin B1 ELISA Test Kit has the Aflatoxin B1acity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8℃ *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Aflatoxin B1 Plate | 1x 96-well Plate (8 wells x 12 strips) | 2-8℃ 2-8℃ |
| Blank Transfer Plate | 1x 96-well Plate (8 wells x 12 strips) |  |
| Aflatoxin B1 Standards: Negative control (white cap tube) 0.05ng/mL (yellow cap tube) 0.25ng/mL (orange cap tube) 0.75 ng/mL (pink cap tube) 2.5ng/mL (purple cap tube) 10 ng/mL (blue cap tube) 1000 ng/mL (spiking, red cap tube) | 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL |  |
| Aflatoxin B1 HRP | 10 ml |  |
| 20X Wash Solution | 28mL |  |
| Stop Buffer | 14 mL |  |
| TMB Substrate | 12 mL |  |
| 10XSample Extraction Buffer | 25ml |  |

* 	If you are not planning to use the kit for over 1 month, store Aflatoxin B1 Standards and Aflatoxin B1-HRP Conjugate at -20°C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Milk/Milk Powder | 0.1 |
| Dried Meat/Fish, Cereal, Feed, Seed | 1 |
| Beef, Pork, Poultry | 1 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Aflatoxin B1 | 100.0 |
| Aflatoxin B2 | 23 |
| Aflatoxin G1 | 36 |
| Aflatoxin G2 | 22 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- 100%Methanol

### Warnings and Precautions

- The standards contain Aflatoxin B1 . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25℃ / 68 – 77°F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25℃ / 68 – 77°F) or in a refrigerator before use.

- Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

- Preparation of Solution A used for Milk:

Mix 6 volumes of 100% Methanol with 4 volumes of 1X Sample Extraction Buffer.

-  Preparation of Solution B used for Dried Meat/Fish, Cereal, Feed, Seed:

Mix 6 volumes of 100% Methanol with 14 volumes of 1X Sample Extraction Buffer.

-  Preparation of Solution C used for Meat: Beef, Pork, Poultry:

Mix 7 volumes of 100% Methanol with 23 volumes of 1X Sample Extraction Buffer.

-  Preparation of 70% Methanol:

Mix 7 volumes of 100% Methanol with 3 volumes of distilled water.

-  Preparation of 87.5% Methanol:

Mix 7 volumes of 100% Methanol with 1 volume of distilled water.

When preparing solutions, vortex vigorously for at least one minute in an air-tight container.

Milk

- To 1 mL of milk, add 1 mL of Solution A.

- Vortex vigorously for 1 minute.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25℃ / 68 – 77°F).

- Use1 50 µL of the sample per well for the assay, avoiding any fat that may have formed on top of the sample.

Note: Dilution factor = 2

Milk Powder

- To 1 g of milk powder in a vial, add 9 mL of distilled water. Vortex vigorously for 2 minutes to obtain a homogeneous suspension.

- Follow the milk sample preparation protocol outlined above, using 150 µL of the sample per well for the assay, avoiding any fat that may have formed on top of the sample.

Note: Dilution factor = 2

Dried Meat/Fish, Cereal, Feed, Seed

- Grind and mix a representative sample (according to accepted sampling techniques).

- Weigh out 5 g of the ground sample and place into a suitable container.

- Add 25 mL of 70% Methanol* and vortex vigorously for 3 minutes manually or 10 minutes using a multi-vortexer.

- Centrifuge the samples for 10 minutes at 4,000 x g at room temperature (20 – 25℃ / 68 – 77°F).

- Dilute 100 µL of the obtained supernatant with 700 µL of Solution B. Vortex for 1 minute.

- Use 150 µL of the diluted sample per well in the test.

Note: Dilution Factor = 40

*The sample size may be increased if required, but the volume of methanol/water must be changed accordingly, for example, if a 10 g sample is to be used, 50 mL of 70% methanol should be used. If the Aflatoxin B1 concentration is expected to be high, the sample can be further diluted with 35% Methanol in 1X Sample Extraction Buffer.

Beef, Pork, Poultry

- Remove fat from a representative sample and homogenize the sample using a suitable blender or mixer.

- To 2 grams of a homegenized sample, add 8 mL 87.5% Methanol and vortex vigorously for 3 minutes manually or 10 minutes using a multi-vortexer.

- Centrifuge the samples for 10 minutes at 4,000 x g at room temperature (20 – 25℃ / 68 – 77°F).

- Dilute 300 µL of the obtained supernatant with 900 µL of Solution C. Vortex for 1 minute.

- Use 150 µL of the diluted sample per well in the test.

Note: Dilution Factor = 20

AFLATOXIN B1 ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25℃ / 68 – 77°F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| B1- HRP | 75 L | 1.8mL |
| 1X Wash Solution | 1.0 mL | 24mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

    - Add 150L of each Aflatoxin B1 Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

    - Add 150L of each sample in duplicate into different sample wells.

    - Add 75 L B1-HRP to each well and mmix well by gently rocking the plate manually for 5 times.

    - Transfer 150  µL of the mixture created in Step 3 from the transfer plate wells into the Aflatoxin B1 plate wells.It is highly recommended to use a multi-channel pipette for this step

    - Incubate the plate for 10 minutes at room temperature (20 – 25℃ / 68 – 77°F).

    - Wash the plate 3 times with 250L of 1X Wash Solution.After the last wash, invert the plate and gently tap the plate dry on paper towels(Perform the next step immediately after plate washings.Do not allow the plate to air dry between working steps).

    - Add 100L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubation for 15 minutes in room temperature, add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Aflatoxin B1 Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Aflatoxin B1 standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us